## Slide 1: FORM-3D: Fast Optimization of Range Modulators 3D Grant Giovani INFN, November, 19th, 2024
Alberto Taffelli
Rome, November, 19th, 2024

## Slide 2: INTRODUCTION
Grant Giovani INFN, Roma, Nov 19th, 2024
A. Taffelli, PhD
Schuller et al., 2020
Dose (a.u.)
Proton Therapy:
Protons can deliver dose precisely in deep seated tumours thanks to their peaked dose distribution in depth (Bragg peak), thus sparing normal tissues
Dell’Oro et al., 2021
FLASH Effect:
Ultra-high dose rates (UHDR) irradiation, i.e. >40 Gy/s, have a protective effect on normal tissues, compared to conventional irradiation (0.01Gy/s), while maintaining the same efficacy against tumours.
In RT = increase the therapeutic window
Active Pencil beam scanning (PBS) systems make use of different energy layers to reproduce dose uniformity in the target (SOPB)
SOBP
n ∽ 103-104
COST FUNCTION

## Slide 3: INTRODUCTION
FLASH Effect in proton therapy:
Problem: modern PBS proton therapy centers (PTC) can’t deliver UHDR due to the energy switching time required between different energy layers (0.5-2 s)
The FLASH effect is lost
3D Range Modulators (3DRM)
Solution: 3DRM, objects able to passively modulate the proton pencil beam energy and conformally shape the dose in the target (tumor)
Optimization based on pre-calculated look-up-tables
Each pin is designed to contribute correctly to the final dose distribution (pin’s lengths)
Fluence of each pencil beam needs to be optimized
Grant Giovani INFN, Roma, Nov 19th, 2024
A. Taffelli, PhD
Simoneov et al. 2017

## Slide 4: PURPOSE
Why?
The design of 3DRM is a non-trivial optimization process, dealing with inaccurate 3D look-up-tables, requiring long computational times, uncompatible with clinical routine (up to days)
A complete workflow to design, print and test the 3DRM is still missing
To speed up the translation of FLASH proton therapy into clinics
How?
Exploiting the superior computational speed of FRED, a GPU-based MC simulation tool, in the optimization process
Creating a dedicated efficient optimization algorithm to design the 3DRM
Developing a complete workflow for the design, printing and testing of patient-specific 3DRM at the PTC of Trento (APSS)
Fast Optimization of Range Modulators 3D
FRED EXPERTISE:
FRED Commissioning during FRIDA:
APSS beam model implemented in FRED
Tool to recalculate patient’s RT plan developed
Validated on a set of significant patients
www.fred-mc.org
A. Taffelli, PhD
Copyright © 2015-2023. A. Schiavi and Sapienza university of Rome. All rights reserved
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 5: 3DRM DESIGN
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 6
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
TPS dicom files:
RT plan
RT dose
CT
Structures
pb: 001 1; E=105.6; pb: 001 1; E=105.6;
pb: 002 1; E=105.6; pb: 002 1; E=105.6;
… …
pb: 301 1; E=98.5; pb: 301 1; E=105.6;
pb: 302 1; E=98.5; pb: 302 1; E=105.6;
… …
pb: 701 1; E=78.3; pb: 701 1; E=105.6;
pb: 702 1; E=78.3; pb: 702 1; E=105.6;
… …
RT original plan
3DRM based plan
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 7: 3DRM DESIGN
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
Quick guess of the 3DRM based on:
Geometrical considerations on the target
Pre-optimized independent 2D SOPBs
cm
Raw 3DRM
Target structure (e.g. PTV)
Depth (cm)
TPS dose distribution
Depth (cm)
Gy
A. Taffelli, PhD
Starting from simple geometries (e.g. sphere)
Grant Giovani INFN, Roma, Nov 19th, 2024
Gy

## Slide 8: 3DRM DESIGN
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
Dose (Gy)
Depth (cm)
cm
RM #vox | Phantom #vox | nPrim | FRED Time
106 | 106 | 105 | 2 s
106 | 106 | 106 | 2 s
106 | 106 | 107 | 4 s
6.4 x 107 | 106 | 105 | 3 s
6.4 x 107 | 106 | 106 | 3 s
6.4 x 107 | 106 | 107 | 7 s
109 | 106 | 105 | 20 s
109 | 106 | 106 | 15 s
109 | 106 | 107 | 23 s
FRED simulation
Time estimate
A. Taffelli, PhD
p+ beam
Grant Giovani INFN, Roma, Nov 19th, 2024
Very fast recalculation!

## Slide 9
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
dose influence matrix
optimize the weights wj to converge to the objective dose distribution Di
e.g. (PTV coverage, OARs sparing)
dose in the ith voxel
weight of the jth pb (fluence)
Fluence Optimization
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 10
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
Compare DVH metrics: D98%, D2%, Dmax
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
Deffet et al., 2023

## Slide 11
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
STEP 0
STEP 1
STEP 2
Development of an algorithm to correct the pin’s weights based on the dose distribution
(e.g. matching algorithm)
3DRM DESIGN
Time estimate:
(10s (FRED) + 1min (fluence opt)) x 100 iter =
≈ 1-2 hours
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
Fluence
Opt
Fluence
Opt
+
+

## Slide 12
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 13
PRINTING
3D printing performed with a stereolithographic printer in Vero resin (StratasysTM) at UniTN-DICAM
TESTING
Perfomed at the Proton Therapy Center of Trento (APSS):
Experimental room
  To test first 2DRM prototypes
Gantry
  Through the fabrication of a dedicated support and alignment system for the 3DRM (ProMTM)
  with dosimetric Tests and QA
R. Cristoforetti (M. Thesis), 2022
PRINTING & TESTING
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 14
OUTLOOK
A. Taffelli, PhD
FORM-3D aims to make a concrete step towards the translation of FLASH proton therapy into clincs, in the framework of the PTC of Trento, to be possibly extended to other facilities
(but not only)
Grant Giovani INFN, Roma, Nov 19th, 2024
A part from Flash Radiotherapy, the project will be important also:
To speed up the current PT treatments, with an impact in the treatment of moving targets (e.g. Lung) and to better allocate the PTC resources
Create the basis for the use of FRED as an independent MC tool at the PTC of Trento for further research
Fast calculation of dose rate metrics with biophysical implications
Contribute to the development of FRED, due to a ongoing collaboration
Possible technological transfer

## Slide 15
COSTS
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 16
GANTT & WORKPACKAGES
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 17
ORGANIZATION
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
Primo Ricercatore

## Slide 18: FORM-3D: Fast Optimization of Range Modulators 3D Grant Giovani INFN, November, 19th, 2024
Alberto Taffelli
Rome, November, 19th, 2024

## Slide 19
BACKUP SLIDES
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024

## Slide 20
1. Acquire and manage TPS data
2. Initial guess of the shape of a 3DRM
3. FRED simulation with the designed 3DRM
4. Fluence optimization
5. Check clinical goals
PRINTING PROCESS
correction of the weights wl associated to the pin’s lengths
ITERATIVE LOOP
NO
YES
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
1. Acquire and manage TPS data
2. Computation of 3D look-up-tables (LUTs)
3. Optimization of Pin lengths and fluences
4. MC Simulation
5. Check clinical goals
PRINTING PROCESS
YES

## Slide 21
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
Ma et al. 2023

## Slide 22
3DRM DESIGN
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024
Deffet et al. 2023

## Slide 23
DADR PBDR metrics
A. Taffelli, PhD
Grant Giovani INFN, Roma, Nov 19th, 2024