## Slide 1
Proposer: L.Margotti
Perovskite-Hybrid Architecture for Spectroscopy and Energy-Counting of γ-Rays
Members: Dr. A. Ciavatti, Dr. I. Fratelli, Dr. S.M. Carturan, Dr. C. Sbarra, Prof. B. Fraboni
CSN 5 - GRANT GIOVANI RICERCATORI 2025

## Slide 2
PEROVSKITE-Based γ-DETECTORS
A Proof of Concept towards printable, high-performance devices
Gamma SPECTROSCOPY (GS)
Single Pixel
PHOTON COUNTING Dev (PCD)
Single Pixel
1

## Slide 3
PEROVSKITE-Based γ-DETECTORS
A Proof of Concept towards printable, high-performance devices
Gamma SPECTROSCOPY (GS)
Single Pixel
PHOTON COUNTING Dev (PCD)
Single Pixel
2

## Slide 4
PEROVSKITE-Based γ-DETECTORS
A Proof of Concept towards printable, high-performance devices
Gamma SPECTROSCOPY (GS)
Single Pixel
PHOTON COUNTING Dev (PCD)
Single Pixel
Detector type: DIRECT
Area: A < 1mm2
Target aspect ratio: (10:1)
Key feature: high thickness for enhanced interaction volume
Detector type: INDIRECT/HYBRID
Area: A < 1mm2
Target aspect ratio: (1:1)
Key feature: organic-plastic scintillator on top for UV emission
3

## Slide 5
SINGLE CRYSTALS VS THIN FILMS
State of the art: coupling bulk performance with handling features
400
FAPbBr3 – 3mm
γ event
241Am (60keV)
MAPbI3 - 110 µm
Li, M., Wang, S., Wood, A. et al.
Nat. Mater. (2025)
Sakhatskyi, K. et al.
Nat. Photon. (2023)
PEA2PbBr4
Basiricò, L. et al.
Adv. Health. Mat. (2024)
PEA2PbBr4
Margotti, L. et al.
IEEE NSC, MIC, RTSD 2024
300V
α event from 241Am (5.48MeV)
4

## Slide 6
SINGLE CRYSTALS VS THIN FILMS
Coupling bulk performance with handling features
PHASE-g
Device GOALS
LOW COST
SCALABLE
PIXEL DEVICES FOR γ-COUNTING/ ENERGY SPECTROSCOPY
IDEA
3D-ELECTRODE ARCHITECTURE based on SOLUTION PROCESSED 2D-PEROVSKITE
Performance TARGETS
SINGLE-DIGIT ENERGY RESOLUTION
(<2 pcs/mm2) DARK COUNT RATE
FOR CLINICALLY RELEVANT ENERGIES
5

## Slide 7
PEROVSKITE-Based DETECTORS
How do we reach bulk-like performance?
Solution processing is a must!
6

## Slide 8
PEROVSKITE-Based DETECTORS
How do we reach bulk-like performance?
Envisioning PCB-printed Multielectrode Arrays…
Commercially available PCB
7

## Slide 9
Features of interest compared to 3D perovskites
✅ Larger bandgap → Fewer thermally excited carriers ✅ Lower trap density → Reduced trap-assisted leakage ✅ Dielectric confinement → Less unintentional carrier injection
✅ Reduced ion migration → Limited degradation
Lower dark current (10-13A @ 1kV/mm)
Higher stability
2D PEROVSKITES: PEA2Pb(Br/I)4
Featuring highly resistive, more stable structures
Quantum + dielectric confinements
8

## Slide 10
ORGANIC-PLASTIC SCINTILLATOR
Materials and Coupling
PMPS :polydimethyl-co-diphenylsiloxane scintillators
Phenyl groups /dyes: examples
Goals:
✅Peak brilliance in the UV range (320-380nm) ✅Light yield under γ-sources at RT (60% rel. EJ212 under
137Cs γ-rays, i.e. 5/7x103 ph/MeV))
✅Precise optical coupling with 2D perovskites
M. Dalla Palma et al. , 2024,
IEEE Transactions on Nuclear Science
9
Tunable properties through chemical optimization!

## Slide 11
OPTIMIZATION OF GEOMETRICAL PARAMETERS
GEANT4 for feedback-driven analyses
PARAMETERS OF INTEREST
Electric field inside the trenches
Structural features (thickness, width, area)
Optical coupling (emission and absorption matching)
FEEDBACK-DRIVEN SIMULATIONS will provide optimal combination of parameters according to fabrication and desired outcomes.
Required material physical parameters (density, refractive index, charge carriers mobility, μτ product, absorption/transmission coefficients, light yield) have been assessed in previous works or will be determined through dedicated tests onsite.
10

## Slide 12
CORE CHARACTERIZATION
Device-level assessments of GS and PCD
  1. STRUCTURAL
  ANALYSIS
  2. BASIC PERFORMANCE ASSESSMENTS
  3. RADIATION DETECTION
Cross-section (SEM) for perovskite filling
Surface topography (AFM / 3D profilometry) for excess material.
X-rays: sensitivity, time response, dose-rate linearity, LOD, radiation hardness.
α-particles: single pulse detection, μτ product.
J-V and C-V curves, PL and Abs spectra
Pulsed analyses
Scintillator photon yield and spectral distribution.
Ledee et al.
Adv. Optical Mater. 2022
PL spectrum of PEA2PbBr4 thin film
11

## Slide 13
TESTING under GAMMA SOURCES
Performance assessments
PC: FAST CSA (0.16 fC ENC) + 3.2 GS/s DIGITIZER for nanosecond-scale pulse acquisition.
GS: low-noise CSA + shaper (up to 100μs) + Digital MCA
Single-photon detection capability
Sensitivity
Count-rate stability
Impact of hybrid contribution
  Charge Collection Efficiency
  Linearity
  Energy resolution
Available Gamma Sources |  | 
LOWEST ENERGIES | HIGHEST ENERGIES | MID-RANGE ENERGIES
241Am | 137Cs - 60Co | 133Ba
59.5keV | 662keV – 1.17 & 1.33MeV | 81, 160, 223, 356 keV
Lines in clinically relevant ranges for SPECT (99mTc, 140keV)!
Tuneable Source-Sample distance
(1-50 mm)
12

## Slide 14
APPLICATIONS AND PROSPECTS
Expected impact
Compound detectors for nuclear physics
✅ Tiled arrays for localized γ-radiation mapping ✅ Low-dose γ-monitors
Onboard dosimeters
✅ Low-dose γ-monitors
Potential follow-up for the ASI-VUS3 IRIS project (INFN-TTlab) for Wearable Dosimeters for the Crew Members Onboard the ISS.
New prospects for γ-cameras:
✅ Reduction in volume and cost ✅ Extension to printable and large area devices ✅ Real-time monitoring of radiotracers and spectroscopic dosimetry
Sinergy with the POR-FESR FORTRESS PROJECT (Fraboni) at IRCSS Reggio Emilia, which is now provided a novel SPECT-CT (source: 99mTc, major γ-peak at 140keV).
13

## Slide 15
TECHNICAL CONSIDERATIONS
Risk Management and Requests
14
VOICE | YEAR 1 | YEAR 2
Missions | 5k | 5k
Disposables | 10k | 5k
SMU 4200A-SCS* | 53k | 
Pulsed Meas. U. |  | 37k
DMCA + CSA |  | 13k
SPC-Module |  | 5k
TOTAL | 68k | 65k
COSTS
* Co-funding with Dip.Fis. Bologna
Possible Issues and Mitigation
Empty trenches / excess material
Sonication-assisted deposition / lapping techniques
Environmental issues / external light contribution
Deposition of Parylene-C encapsulation layer
Poor γ absorption
SJI / Metal Ink deposition for trenches height increased up to 500μm.
Interface losses between perovskite and scintillator
Dedicated resin / cement of matched refractive index
Delays in electronic instrumentation delivery
Analog spectroscopic chain / Electrometer …

## Slide 16
Conclusions
PHASE-g aims at validating a novel approach to perovskite-based γ-ray detectors, which merges the bulk properties of Single Crystals with the major advantages of interdigitated structures.
It provides:
➡ An Innovative dual structure which relies on 3D interdigitated electrodes
➡ A dedicated study on PMPS-based scintillating device for photon counting
➡ A spectroscopic device for γ-rays of different doses and energies
➡ Tests with clinically relevant radiation energies
It opens up new paradigms for detectors fabrication and thus scalability to large-area, pixelated multielectrode arrays.